## Slide 1
VISHWAKARMA AWARDS 2026-27 · STAGE 2 DESIGN SPECIFICATION
Vishwakarma Application ID: VK26-IITI-HT-0261
VisionAid
A chest-worn, fully offline navigation aid that adds head-level-obstacle and drop-off warnings to the white cane - with an FPGA "safety island" whose alerts keep working even if the AI computer fails.
Team ID: VK26-IITI-HT-0261
Track: Assistive Solutions & Inclusive Living
Team SVNIT: Jayant Kumawat (lead), Sneha Ahir, Aman Gupta - SVNIT Surat
Mentor institute: IIT Indore
Design files: github.com/silentaf/viswa_id1
CAD v1 of the chest pod (FreeCAD) - design stage, not yet fabricated

## Slide 2: Problem Context
43.3 M
people blind worldwide (2020)
Cited: Bourne et al., Lancet GH 2021
295 M
moderate / severe vision impairment
Cited: Bourne et al., Lancet GH 2021
1.99 %
blindness among Indians aged 50+
Cited: NBVIS India 2015-19
₹3,500
SmartCane - closest Indian aid
Cited: IIT Delhi / Assistech
Problem statement
The white cane only senses what it touches at ground level, about one cane-length ahead.
Head-level obstacles (branches, poles, signs, truck bodies) and drop-offs (open drains, kerbs, steps) are what it misses. In a survey of 307 blind / legally blind travellers, 13 % hit their head at least once a month and 7 % fell at least once a month; 23 % of head-level accidents needed medical care, and cane vs guide-dog users did not differ (cited: Manduchi & Kurniawan, 2011; US-based sample).
Reading a label or sign still needs a sighted helper or a hand-held phone.
Target beneficiary: blind and low-vision cane users who walk independently on Indian streets and campuses.
Public search disclosure - what exists
Product | Price (cited) | Gap
SmartCane (IIT Delhi) | ≈ ₹3,500 (2016) | Above-knee only; no drop-offs, no naming / reading
OrCam MyEye 3 Pro | ≈ US$4,250 | Reading aid, not mobility; expensive
Envision Glasses | US$1,899-3,499 | Reading / description; expensive
Phone apps (Lookout, MANI) | free | Need a hand + pointed phone; no continuous warning
Core innovation: head-level + drop-off warnings for the cane user, with an alert path bounded in hardware (FPGA) and an on-device AI layer that names hazards and reads text - fully offline.

## Slide 3: User Journey - Before (cane only)
Walks with cane
cane sweeps the ground ~1 m ahead
Head-level obstacle
not touched by the cane
Collision
head / face injury risk
Open drain or step down
found only when the cane drops in
Needs to read a label
asks a stranger / holds a phone
Current pain points
Warning comes from contact: the cane reports an obstacle only when it touches it.
Nothing above the cane: overhanging objects are invisible to it.
Drop-offs are found by falling-in of the cane tip - the user is already at the edge.
Reading and identifying objects depend on another person.
Quantified inefficiency (calculated, not measured)
Warning time for ground obstacles: ~1 m cane reach ÷ 1.2 m/s walking speed ≈ 0.8 s.
Warning time for head-level obstacles: 0 s - the cane never touches them.
Warning before a drop-off: only as far as the cane tip; no advance cue.
Text: every label / sign needs a helper; no independent reading.
Assumptions: 1.2 m/s typical adult walking speed and ~1 m cane reach are design assumptions; Stage 3 user interviews and trials will replace them with measured values.

## Slide 4: User Journey - After (cane + VisionAid)
Walks with cane + pod
pod scans 0.3-3 m, waist to head
Obstacle in 3 m
FPGA classifies distance zone
L / R vibration
pulse rate = distance; side = direction
Floor drops ahead
LiDAR: drop-off alarm + buzzer
Presses READ
Pi reads text aloud (offline)
How the user benefits
Warned before contact - including head-level obstacles the cane cannot reach.
Left / right motors on the chest straps say which side the hazard is on.
Drop-off warning ahead of the edge, not at it.
Hands stay free; speech through a bone-conduction headset keeps the ears open to traffic.
Alerts never depend on the AI computer (FPGA safety island).
Quantified efficiency (labels say where each number comes from)
Warning time, 3 m range at 1.2 m/s: ≈ 2.5 s (calculated, target range)
Echo edge → motor on: max 214 ns over 4 events (simulated; target ≤ 1 ms)
Update per direction: 99 ms (by design, simulated)
Drop-off warned: 1.86 m before the step (calculated from CAD angles); alert 22 ms after (simulated)
Image → speech audio ready: 211 ms (measured on laptop, not Pi 5; excludes capture + playback)
Honesty rule: no hardware built yet - simulated = our Verilog testbench; measured = our Pi software run on the dev laptop; Pi 5 and field numbers come in Stage 3.

## Slide 5: System Architecture - Block Diagram
3 × RCWL-1601 ultrasonic
(left, right, head-level)
TF-Luna LiDAR
(floor, −35°)
3 buttons
(READ / MODE / QUIET)
FPGA SAFETY ISLAND
Sipeed Tang Nano 9K (27 MHz)
echo timing · distance zones
LiDAR drop-off · PWM patterns
watchdog · fault detection
Left / right coin motors
(chest straps, wired)
Buzzer
("check device", drop-off)
AI PATH - Raspberry Pi 5
YOLO-nano object names · OCR
text-to-speech (EN / HI)
Camera Module 3
(autofocus)
Bluetooth bone-conduction
headset (speech)
UART 115200 status / config
+ heartbeat + fault line
POWER
BIS-certified
USB-C power bank
→ Pi 5
→ 500 mA PTC
→ carrier 5 V
→ AMS1117 3.3 V
Safety rule: the FPGA drives the motors itself; the Pi can only add information and tune thresholds within clamped limits (it cannot switch alerts off). If the Pi's heartbeat stops, alerts continue and an "AI offline" chirp sounds.

## Slide 6: Target Specifications & Hardware Selection
Parameter | Target | Evidence so far | Selected hardware / protocol
Obstacle range (waist-head) | 0.3-3.0 m | sensor class; housings +25° / ±15° (CAD) | 3 × RCWL-1601, 3.3 V, sequenced
Alert latency (echo → motor) | ≤ 1 ms | Simulated: max 214 ns (4 events) | Tang Nano 9K GW1NR-9 FPGA, 27 MHz
Update period / direction | ≤ 100 ms | 99 ms by design (simulated) | 3 × 33 ms slots (no crosstalk)
Distance resolution | ≤ ±5 cm to 2 m | Timer: 1500 mm in → 1500 mm out, exact (sensor error: Stage 3) | 1 µs echo timer (0.17 mm/count)
Drop-off warning | ≥ 1.0 m ahead | Calculated 1.86 m (−35° from 1.3 m); alert 22 ms (sim.) | Benewake TF-Luna + IMU tilt correction
Object name / OCR | ≤ 1 s / ≤ 3 s | Laptop: detect 28 ms, e2e 211 ms, OCR 1.2-1.4 s | Raspberry Pi 5 4 GB + Camera Module 3
Pi fails → alert path | unaffected, cue ≤ 1 s | Simulated: AI-offline at 666 ms | FPGA heartbeat watchdog (GPIO17)
Runtime | ≥ 4 h | Calc. (cited Pi 5 W): 4.3 h if AI busy ~50 % (assumed); 3.1 h if always on | 10,000 mAh BIS power bank, USB-C
Pod mass | ≤ 250 g | Estimate: likely over (plastic ≈156 g solid) | PETG, 2.5 mm wall; lightening in v2
Architecture rationale: Pi 5 (4 × Cortex-A76, 2.4 GHz) instead of the Stage 1 PYNQ-Z2 (2 × Cortex-A9, 650 MHz) for detection + OCR; FPGA kept for the parallel, cycle-exact, fault-isolated alert path; wired haptics so no radio sits in the safety path.

## Slide 7: CAD Design & Mechanical Specs
FREECAD · PARAMETRIC
Exploded: shell | Pi 5 | 20 mm standoffs | carrier PCB + Tang Nano | lid + housings | sensors
Back: GoPro-style mount, vents
Front: sensor housings
Mechanical specification
Envelope: 112 × 84 × 48 mm shell + front housings
Model: 19-parameter spreadsheet drives the model; shell sketch fully constrained
Housings: head sensor +25°, side sensors ±15° yaw, LiDAR −35°
Mount: GoPro-style 2-finger (3 mm, M5) on a chest harness
Cooling: side + top/bottom vents for the Pi 5 active cooler
Interference check: 0 mm³ vs Pi, carrier, Tang Nano, sensors
Design iteration: adding the Tang Nano 3D model showed a 182 mm³ clash with the lid (171 mm³ of it the HDMI connector) - shell depth raised 44 → 48 mm in one parameter; clash now 0 mm³.

## Slide 8: Material Selection & FEA (chest mount)
SIMULATED
v1 sharp roots - 20 N side
v2 1.5 mm fillets - 20 N side
Material | Tg / softening | Choice
PLA | ≈ 60 °C | No - softens in a car / summer sun
PETG | ≈ 80 °C | Yes - tough, easy to print
ASA | ≈ 100 °C, UV-stable | Stage 3 option outdoors
FreeCAD FEM + Gmsh + CalculiX
Load A: 50 N down (0.3 kg × ~17 g bump); B: 20 N sideways
PETG: E 2.0 GPa, ν 0.38, yield ≈ 50 MPa; required FoS ≥ 3
Result: v2 worst FoS 8.7 (every run ≥ 8.6); max deflection 0.064 mm
v1: root stress kept rising with mesh refinement → sharp-corner singularity
v2: 1.5 mm root fillets → converged (≤ 3 % spread, 3 meshes); now in the CAD
Limitations: linear isotropic model (printed parts are anisotropic); finest v2 mesh failed to solve, so v2 uses 3 mesh levels; a printed-mount pull test is planned for Stage 3.

## Slide 9: Electronics - Schematic & Carrier PCB
KICAD 9
Carrier-board schematic (A3): power, Pi header, FPGA, sensors, haptics, buttons, test points
Power distribution & checks
Power bank → Pi 5 → header 5 V → 500 mA PTC → +5 V
+5 V → AMS1117 → +3V3 (sensors, motors, IMU)
+5 V → SS14 → Tang Nano (blocks USB back-feed)
Verified: ERC 0 · DRC 0 · unrouted 0 · parity 0
PCB: 85 × 56 mm, 2 layers, Pi 5 outline + mounting holes, 20 mm standoffs + 40-way ribbon; 3.3 V-bank FPGA pins only; MOSFET motor drivers with flyback diodes; Gerbers generated.

## Slide 10: Bill of Materials & Sourcing
Component | Part / specification | Qty | Source · origin · lead | ₹ (basis)
AI computer | Raspberry Pi 5 4 GB + active cooler | 1 | KSP Electronics · imported · out of stock | 15,050 (quote + est.)
FPGA | Sipeed Tang Nano 9K | 1 | Probots · imported · out of stock | 2,999 (quote)
Camera | Pi Camera Module 3 + cable | 1 | Thingbits · imported · in stock | 3,358 (quote + est.)
LiDAR | Benewake TF-Luna | 1 | Probots · imported · out of stock | 2,118 (quote)
Ultrasonic + IMU | RCWL-1601 × 4, MPU-6050 | 5 | Robu · local · 2-4 d | 800 (est.)
Feedback | Coin motors × 2, BT bone-conduction headset | 3 | Robu / Amazon.in · 2-5 d | 2,120 (est.)
Power | 10,000 mAh BIS power bank | 1 | local · 1-3 d | 1,200 (est.)
Carrier PCB | 2-layer fab ×5 + SMD parts + connectors + ribbon | 1 | LionCircuits (domestic) · 5-7 d | 3,250 (est.)
Mechanical | PETG, standoffs, inserts, harness, buttons | 1 | Robu / Amazon.in · 2-5 d | 2,350 (est.)
microSD | 64 GB A2 | 1 | local · 1-3 d | 700 (est.)
TOTAL |  |  | Grant ₹30,000 | 33,945
Budget - honest status
Everything except the Pi 5 board: ₹19,445.
Pi 5 4 GB on 9 Oct: ₹12,600-14,500, out of stock → ₹2,045-3,945 over the grant.
Rule: Pi 5 must be ≤ ₹10,555 to fit; else re-quote or cover the gap from team funds.
Battery life (calculated, cited Pi 5 W)
Pi 5: ≈ 3.0 W idle, ≈ 8.8 W full load (published); rest 1.34 W (est.)
AI always on: 10.1 W → 3.1 h
Event-triggered, ~50 % (assumed): 7.2 W → 4.3 h
AI idle: 4.3 W → 7.2 h
Procurement status: nothing ordered in Stage 2. Quotes are live listings checked 9 Oct 2026; Pi 5, Tang Nano 9K and TF-Luna were out of stock that day - re-quote and order early. Itemised BOM: design/bom/VisionAid_BOM_Power.xlsx.

## Slide 11: FPGA Safety Island - Verification & Implementation
SIM + REAL P&R
Plotted event: 136 ns from echo edge to motor ON (motor rises on the same clock as the zone change). The testbench monitor samples 2 clocks later, so its 214 ns is a conservative upper bound.
Testbench check (Icarus Verilog) | Result
Distance 1500 / 2500 mm | 1500 / 2500 mm (exact)
Echo edge → motor ON | 136 ns wave; ≤ 214 ns bound
Frames to Pi (Verilog / Py) | 192 / 192 decoded, 0 bad
LiDAR drop-off | alert 22.3 ms after change
Pi heartbeat lost | AI-offline 666 ms; alerts on
Sensor unplugged | fault flag + FAULT_N low
Overall | ALL TESTS PASSED
27 MHz | GW1NR-9 nextpnr | Zynq-7020 Vivado
LUTs | 1,561 / 8,640 (18 %) | 658 / 53,200 (1.24 %)
Flip-flops | 574 / 6,480 (8 %) | 574 / 106,400 (0.54 %)
ALU / DSP | 930 ALU (14 %) | 1 DSP (0.45 %)
Block RAM | 0 / 26 | 0 / 140
I/O | 26 / 276 | n/a (OOC run)
Timing | Fmax 79.4 MHz, met | WNS +29.1 ns, met
Routed results (GW1NR-9 bitstream built; Vivado out-of-context). No hardware yet.
Cross-verified: the Verilog FPGA accepted a command built by our Python Pi code, and the Python decoder read every byte the FPGA sent - both sides of the link agree.

## Slide 12: AI Path (Raspberry Pi software) - Proof of Concept
MEASURED · LAPTOP
What runs (same code for the Pi 5)
YOLOv5n (ONNX) → hazard + side; RapidOCR; Piper TTS English + Hindi; alert manager; IMU floor baseline
Event-triggered: AI runs only when the FPGA sees an obstacle < 3 m
Full-loop replay (FPGA sim bytes → speech)
"Step down ahead. Stop." · "Check device: head sensor."
"person, left, 1.6 metres" - WAV files in design/pi/results/replay
Honest scope: speeds measured on an i7 laptop (4 threads), NOT a Pi 5. Accuracy is proof-of-concept, not production-grade: mAP50 0.41 on stock COCO weights (200 images vs human labels). Stage 3: fine-tune on our own Indian, chest-height data - same plan as the pothole model.

## Slide 13: Model Training - Pothole Detector (fine-tuned)
TRAINED · HELD-OUT TEST
16 held-out INDIAN road test photos (RDD2020) - boxes = its predictions, number = confidence
Held-out test split (330 images, 736 potholes) | Result
Precision | 0.56
Recall | 0.42
mAP50 / mAP50-95 | 0.45 / 0.19
Speed, laptop CPU, 4 threads (not Pi 5) | 23 ms / image
Before fine-tuning (COCO model) | no pothole class at all
How it was trained
YOLO11n, COCO-pretrained start, 100 epochs, RTX 5060 GPU, 8.8 min
1,239 train images: Ryukijano public set (153 unique photos, CC BY 4.0) + 1,070 REAL INDIAN road photos (RDD2020, Arya et al. 2020, peer-reviewed)
Checked: zero unique-photo overlap between train, valid, test
Honest limits
Finds ~4 in 10 potholes; a bit over half of detections are real
RDD2020 photos are dashcam/vehicle height, not yet our chest-mounted angle
Stage 3: capture our own chest-height photos with the actual pod camera
Why it matters: recall > precision on purpose - missing a pothole risks a fall, a false alarm costs one caution beep. Test set now includes 230 real Indian road photos (RDD2020), not just the original non-Indian set. Files: design/pi/training.

## Slide 14: Sensor Coverage & Drop-off Analysis
CALCULATED + SIMULATED
Drop-off: target met
LiDAR at −35° from 1.3 m sees a 15 cm step 1.86 m ahead (≈ 1.5 s at walking speed)
Found: walking sway
±3° chest sway → 25 false alarms in 20 s; IMU-corrected baseline → 0 (now in the Pi code)
Found: waist-height gap
Under 1 m ahead, 0.3-1.0 m height is not covered → tilt forward sensors −10° in CAD v2
Assumptions: 1.3 m pod height, 15° ultrasonic half-beam (datasheet class), ±3° sway at 1.8 Hz, 1 cm LiDAR noise - design/analysis/sensor_geometry.py.

## Slide 15: Your Maker Journey - Evidence
System data-flow (simplified)
Sensors
3x ultrasonic + LiDAR + camera
FPGA safety island
(Tang Nano 9K) - zone logic
Haptic motors
+ buzzer
(hardware-bounded)
Pi 5 AI path
YOLO + OCR
(event-triggered)
Bluetooth headset
spoken alerts (Hi / En)
Full architecture + timing on slide 5 (System Architecture)
FreeCAD feature tree - FingerRootFillet R 1.5 mm
KiCad DRC: 0 violations, 0 unconnected · ERC: 0 errors, 0 warnings
GTKWave: echo falls → distance 749 mm → zone 3 → left motor ON (simulation)
Raw repository (native files)
github.com/silentaf/viswa_id1 - .FCStd, .step, .kicad_sch/.kicad_pcb, .v, .xlsx, timestamped commits
Stage 1 → 2: PYNQ-Z2 → Pi 5 + Tang Nano 9K; per-stage latency specs; LiDAR for drop-offs; FEA-driven fillet; CAD clash fix 44 → 48 mm (details on the next slide).

## Slide 16: Stage 1 → Stage 2 Engineering Iterations
Stage 1 said | Stage 2 decision | Why
PYNQ-Z2 does everything | Pi 5 (AI) + Tang Nano 9K FPGA (safety island) | 2 × A9 650 MHz vs 4 × A76 2.4 GHz for AI; fault isolation
"<10 ms" reflex, "<100 ms" on cover | Per-stage specs: echo → motor ≤ 1 ms | Echo from 3 m alone takes ≈ 17.5 ms - "<10 ms end-to-end" is impossible
Downward ultrasonic for drop-offs | Downward TF-Luna LiDAR (−35°) | Narrow beam gives a clean floor-distance signal
Wristband with 2 motors | L / R motors on the chest straps, wired | Two motors on one wrist are hard to tell apart; no radio in safety path
Custom bone-conduction driver | Bluetooth bone-conduction headset | Pi 5 has no analog audio; speech is the information channel
"< ₹5,500 BOM" | Prototype BOM ₹33,945 at 9 Oct quotes | Pi 5 + camera alone exceed ₹5,500
"250M+", "₹4-6 lakh" competitors | Cited 2020 data; SmartCane disclosed | Accuracy and honest public-search disclosure
10-15 pilot users | Stage 3, after ethics approval + O&M instructor | Honest scope for a design round
Found by our own analysis: FEA → 1.5 mm root fillets; CAD clash → shell 44 → 48 mm; sway simulation → IMU floor baseline; power analysis → event-triggered AI; tests → OCR word-order bug fixed.

## Slide 17: Safety & Hazard Mitigation
Failure mode | Hazard | Embedded fail-safe / mitigation
Ultrasonic stuck / unplugged | Missed obstacle | No-response detector (3 slots) → buzzer "check device" + FAULT_N (simulated)
Pi crash / hang | No object names / OCR | FPGA watchdog → "AI offline" chirp; alert path unaffected (simulated)
Bluetooth headset drops | No speech | Haptics carry all safety alerts; distinct cue
Walking sway / tilt | False drop-off alarms | IMU-corrected floor baseline: 25 → 0 false alarms (simulated)
Pi 5 overheats in pod | Throttling, slow AI | Active cooler + vents in CAD; thermal test in Stage 3
Carrier short circuit | Pi 5 V rail damage | 500 mA PTC fuse on the header supply
Battery fire | Burns | BIS-certified power bank only - no loose Li-ion cells
Alert fatigue | User ignores alerts | Zone hysteresis + quiet mode (silences "info" only)
Standards considered in the design (not certified)
IEC 62368-1 / IS 13252 - ICT equipment safety
IEC 62133-2 / IS 16046 - Li-ion (via BIS power bank)
IEC 60529 - IP54 target for the pod
ISO 10993-5/-10 - skin-contact materials (harness)
ISO 9999 - assistive-product classification
WPC / ETA - only pre-certified radio modules
Design principle: the cane stays the primary aid - VisionAid only adds warnings, and its alert path is independent of the AI computer.

## Slide 18: Roadmap - 4-Month Fabrication Plan
Nov 2026
Order parts + carrier PCB (LionCircuits)
FPGA + sensor bring-up on the bench
Range, latency (logic analyser at TP4/TP5), power tests
Dec 2026
PCB assembly; Pi software integration
Print + fit enclosure; thermal test in pod
India dataset: rickshaw, cow, drain (pothole done early, RDD2020)
Jan 2027
Ethics approval (institute ethics committee)
Supervised trials: sighted blindfolded first
Then blind volunteers with an O&M instructor
Feb 2027
Fix top issues from trials
Re-run all Stage 2 tests on hardware
Field-test-ready unit for Stage 3
Stage 3 commitment
A wearable, battery-powered VisionAid prototype ready for supervised functional field testing, with every Stage 2 target (range, alert latency, drop-off distance, runtime, mass) measured on hardware.
Research extension: RISC-V soft core (PicoRV32) on the Tang Nano + a tiny spiking neural network to classify sensor patterns (stairs / wall / person) at very low power - once real sensor data exists.

## Slide 19: Technical Challenges & Risks
Ultrasonic crosstalk vs update rate
Three sensors firing together hear each other. Sequencing fixes it but limits each direction to one ping per 99 ms (simulated).
Enclosure clash found in CAD
With the Tang Nano 3D model added, the module clashed with the lid by 182 mm³ - 171 mm³ of it the HDMI connector (CAD interference check).
Weight over target
Enclosure plastic alone ≈ 156 g if solid (CAD volume); with Pi 5 and sensors the pod likely exceeds 250 g.
Budget over the grant
At 9 Oct quotes the BOM is ₹33,945 - ₹3,945 over the ₹30k grant; Pi 5, Tang Nano and TF-Luna were out of stock.
Battery vs always-on AI
With the AI always running, the calculated runtime is 3.1 h - below the 4 h target (cited Pi 5 power).
Sway false alarms + coverage gap
Simulation: 25 false drop-off alarms in 20 s with a fixed floor baseline; waist-height gap under 1 m ahead.
Source of each risk: our own design files - CAD interference report, FEA study, BOM workbook, FPGA simulation - not assumptions copied from elsewhere.

## Slide 20: Potential Solutions & Risk Mitigation
Risk | Mitigation / backup
Crosstalk vs rate | Keep sequencing; Stage 3 option: coded pings to fire in parallel
Enclosure clash | Done: shell depth 44 → 48 mm (one parameter); interference now 0 mm³
Weight | v2: 2.0 mm walls + 20 % infill, merge housings, Pi Compute Module later
Budget | Pi 5 must be ≤ ₹10,555 to fit; re-quote at order time, else team funds cover the gap
Battery | Done in code: event-triggered AI → 4.3 h (calc., ~50 % AI duty assumed); measure on hardware
Sway / gap | Done: IMU floor baseline (0 false alarms, sim.); gap: forward sensors −10° in CAD v2
Mount corners | Done: 1.5 mm root fillets (FEA v2 converged, FoS ≥ 8.7)
Pi failure / heat | FPGA watchdog + independent alert path (simulated); cooler + vents, thermal test
Carrier PCB with the Tang Nano 9K model (KiCad)
Status: "Done" items are already in the design files; the rest are planned and will be verified on hardware in Stage 3.

## Slide 21: References & Sources
References
Bourne R. et al., Trends in prevalence of blindness and vision impairment, Lancet Global Health 9(2):e130-e143, 2021.
National Blindness & Visual Impairment Survey India 2015-19, AIIMS / MoHFW (indiavisionatlasnpcb.aiims.edu).
Manduchi R., Kurniawan S., Mobility-related accidents experienced by people with visual impairment, Insight 4(2), 2011 (tech. report UCSC-SOE-10-24; 307 respondents).
SmartCane - IIT Delhi / Assistech: ₹3,500 (Snapdeal launch, 2016; entrepreneur.com, electronicsforu.com).
OrCam MyEye 3 Pro US$4,250 (nelowvision.com); Envision Glasses US$1,899 / 3,499 (shop.letsenvision.com) - checked 9 Oct 2026.
Sipeed Tang Nano 9K wiki and pin map; Benewake TF-Luna datasheet; Raspberry Pi 5 / Camera Module 3 product briefs.
Raspberry Pi 5 power (≈ 3.0 W idle, ≈ 8.8 W full load): raspberry.tips power comparison, 2026. Part prices: kspelectronics.in, electronifyindia.com, probots.co.in, thingbits.in (9 Oct 2026).
COCO val2017 (cocodataset.org); YOLOv5n, YOLO11n (Ultralytics); RapidOCR / PP-OCR; Piper TTS (rhasspy). Pothole data: Roboflow Universe "Potholes Detection" (project-ssayl), CC BY 4.0, via Hugging Face Ryukijano/Pothole-detection-Yolov8.
Arya D. et al., Global Road Damage Detection: State-of-the-art Solutions, IEEE Big Data, 2020 (RDD2020 India subset, CC BY-SA 4.0: 1,530 real Indian road photos with pothole boxes; huggingface.co/datasets/ShixuanAn/RDD2020).
Tools: KiCad 9, FreeCAD + CalculiX + Gmsh, Icarus Verilog, GTKWave, Yosys, nextpnr + Apicula, Vivado 2026.1, FreeRouting.
Image sources
All diagrams, renders, plots and screenshots in this deck are our own, generated from the project files in github.com/silentaf/viswa_id1 (one exception below):
CAD renders - design/cad (FreeCAD model)
PCB / schematic - design/electronics (KiCad)
FEA plots - design/cad/outputs/fea
Waveforms - design/fpga (testbench); AI / analysis charts - design/pi, design/analysis
Screenshots - docs/screenshots
Exception: the road photos on the pothole slide are RDD2020 India photos (CC BY-SA 4.0, see references); only the boxes are our model's output.
Logos: Vishwakarma Awards / Maker Bhavan Foundation (official template).